## Slide 1: Status of the KATRIN experiment
Nancy Wandkowsky for the KATRIN Collaboration
September 20, 2013
KATRIN overview
STS status
SDS commissioning
Transmission studies
Background studies
Summary and Outlook

## Slide 2: KATRIN overview
source
transport
spectrometer
detector
2013: Spectrometer and Detector Section (SDS) commissioning
2014: Source and
Transport Section (STS) commissioning
KATRIN experimental setup

## Slide 3: STS status
Windowless gaseous tritium source
WGTS Demonstrator
06:00 07:00 08:00 09:00 10:00
4h measurement
KATRIN specification
Temperature TBT [K]
30.35
30.25
30.20
30.30
WGTS demonstrator successfully tested in 2011/12
2013/14: upgrade to full source
magnet tests at CEA Saclay
delivery and commissioning in 2014
S. Grohmann et al, Cryogenics 55, 5 (2013)

## Slide 4: STS status
Differential pumping section
DPS pumping concept successfully tested in 2011
gas flow reduction by 2.5·104 achieved
failure of protection diode → new concept worked out
delivery and commissioning in 2014
S. Lukic et al., Vacuum 86, 8 (2012)

## Slide 5: STS status
Cryogenic pumping section
CPS pumping concept demonstrated
beam tube gold coating, cold tests of magnets successfully passed
reworking of protection diodes
delivery and commissioning in 2014
F. Eichelhardt et al., Fusion Sci. Technol. 54 (2008)

## Slide 6: KATRIN overview
source
transport
spectrometer
detector
2013: Spectrometer and Detector Section (SDS) commissioning
2014: Source and
Transport Section (STS) commissioning
KATRIN experimental setup

## Slide 7: SDS commissioning - objectives
Main Spectrometer
angular
selective
E-Gun
Detector
Section
Primary objectives:
test of individual hardware, software and slow control components
provide ultra high vacuum conditions at the p ≈ 10-11 mbar level
detailed understanding of MAC-E filter transmission properties (E = 18.6 keV with ΔE = 0.93 eV resolution)
detailed knowledge and control of background processes

## Slide 8: SDS commissioning – preparations
May 8, 2012
Spectrometer closed

## Slide 9: SDS commissioning – bake-out
back at vacuum in July 2012: 10-8 mbar (water-dominated)
required pressure: 10-11 mbar to minimize scattering on residual gas
challenge: large volume of 1240 m3
p = 10-11 mbar
Motivation

## Slide 10: SDS commissioning – bake-out
turbomolecular
pump
inner surface: 690 m2
required temperature: 200°C
slow heating so inner electrode
can follow expansion of vessel
Procedure

## Slide 11: SDS commissioning – bake-out
back at vacuum in July 2012: 10-8 mbar
required pressure: 10-11 mbar to minimize scattering on residual gas
(water-dominated)
challenge: large volume of 1240 m3
activation of getter material (NEG St 707)
Getter activation

## Slide 12: SDS commissioning – bake-out
slow heating/cooling (~1°C/h): expansion of vessel and electrode
temperature breakpoints: 200°C – water removal from vessel
300°C – activation of getter material
leak developed &
fixed immediately
back to
10-11 mbar level
Today: 4·10-11 mbar
Chart: Baking cycle (4. - 30.1.2013)
Procedure

## Slide 13: SDS commissioning – bake-out
J. Wolf
Chart: 
fix point
expansion
(video mirrored)

## Slide 14: SDS commissioning - detector
Integration of detector and spectrometer
May 16, 2013: SDS connection established
May 28, 2013: opening gate valve

## Slide 15: SDS commissioning – first light
May 31, 2013
asym. magnetic field
map electrode onto detector
confirmed by simulations

## Slide 16: SDS commissioning – studies
magnetic guiding
electrostatic retardation
B [T]
e
e
e
q [°]
Transmission of ß-electrons
magnetic bottle effect
long storage (h)
Background: stored electrons

## Slide 17: Transmission studies – e-gun
Quasi-monoenergetic, angular
selective electron gun
diagnostic tool: mapping of
analyzing plane
K. Hugenberg et al, Prog. Part. Nucl. Phys. 65, 288 (2010)

## Slide 18: Transmission studies – mapping
Quasi-monoenergetic, angular
selective electron gun
diagnostic tool: mapping of
analyzing plane

## Slide 19: Transmission studies
Quasi-monoenergetic, angular
selective electron gun
diagnostic tool
study of electron transmission
First transmission function
at high voltage (Aug. 14th)
electron rate (kcps)
UHV-Uegun (eV)
successful transmission!
width of 100 meV
PRELIMINARY
e
e
0.5
2.5
1.0
2.0
1.5
3.0
-0.5
-0.6
-0.8
-0.7
-0.9
-1.0
-1.1

## Slide 20: Background studies – radon
Radon sources at the KATRIN main spectrometer
getter material (St 707): 219Rn (t1/2=3.96 s)
vessel surface: 220Rn (t1/2=55.6 s)

## Slide 21: Background studies – radon
Radon decay in sensitive spectrometer volume
alpha-emission: not relevant for KATRIN
accompanying electrons: source of background

## Slide 22: Background studies – radon model
F. Fränkle et al, Astropart. Phys. 35, 128 (2011)
S. Mertens et al, Astropart. Phys. 41, 52 (2013)
N. W. et al, J. Phys. G 40 (2013)
N. W. et al, New J. Phys. 15 (2013)
pre-spectrometer
magnetically stored electrons scatter off residual gas
→ ionization → secondary electrons
high primary energies → hundreds of secondary electrons
UHV (10-11 mbar) → storage time up to several hours

## Slide 23: Background studies – expectations
S. Mertens et al, Astropart. Phys. 41, 52 (2013)
modeling of physics processes: KASSIOPEIA
extrapolations to main spectrometer
expectation: background on the order of 1 count per second
Expectations for the main spectrometer
scenario 1
scenario 2
scenario 3
scenario 4
1
10-1
10-2
10-3
10-4
Background rate (cps)
1

## Slide 24: Background studies – observations
no large Penning traps (previous MAC-E filters: >105 cps initial rate)
highly effective magnetic shielding against cosmic μ-induced background
required rate: 10-2 cps
counter measures needed: electric shielding against μ–induced background,
passive/active reduction of stored electron background
First look at background with full HV
Background rate (cps)
2.5
2
1.5
1
0.5
0
100
200
300
400
time (s)
600
500
10-2
2·10-3
PRELIMINARY

## Slide 25: Background studies – reduction
NEG-Pump
LN2 cooled
baffle
spectrometer
vessel
219Rn
Passive background reduction: principle

## Slide 26: Background studies – reduction
Passive background reduction: measurement results
WARM BAFFLES
COLD BAFFLES
Background rate/unit volume (cps/m3)
detector rings
PRELIMINARY
reduction by factor 3

## Slide 27: Background studies – reduction
Active background reduction: overview
Cyclotron Resonance
All electron energies
Magnetic Pulse
High energies
Electric Dipole
Low energies
breaking of electron storage condition
removal from sensitive volume
S. Mertens et al, JINST 7 P08025 (2012)
N.W., PhD thesis (2013)

## Slide 28: Background studies – reduction
Active background reduction: dipole measurement
Background rate (cps)
1
10-1
10-2
103
102
10
0 20 40 60 80 100 120
time (s)
PRELIMINARY
Short application of electric dipole field
E≈1 V/m → ExB-drift
enhanced bg rate by ignition of
Penning trap between egun & spec

## Slide 29
Summary and Outlook
KATRIN: determination of mν <0.2 eV (90% CL)
STS components delivery and commissioning in 2014
SDS Milestones in 2013:
  Jan.: bake-out of the spectrometer
  May: integration of detector system (SDS)
  May: first light
  Aug.: application of HV
Commissioning: transmission and background studied
End of commissioning: Sept. 23rd
Detailed analysis of collected data
Prepare SDS for tritium measurements
Start of tritium measurements in 2015